New in 2008

JWD1000 Detector type Min Series Optical Power Meter

JWD1000 Min Series Power Meter equips with wide angle detector, replaceable attenuator and various fiber adapters, which can be connected directly with the computer with USB or RS232 communication port. It is an intelligent but economical test device and very suitable for optical power measurements from Passive Components Producers, Research Lab and University Lab.

Various optical fiber adapters and attenuation chips can be used upon actual requirements.

Product Features

- Wide angle Si、Ge detectors can be alternative

The JWD 1000 min series OPM can support the detector size up to 5mm for Ge detector and 10mm for Si detector.

The measurement wavelength range is from 400~1100nm for Si detector and 800~1600 for Ge detector, with dynamic ranger of 60dB.

- Compensation of the detector temperature

JWD1000min series OPM can calibrate the value of the testing result trough the detector temperature compensation; this will lessen the reading value offset caused by the external temperature change

- Various Optical Fiber Adaptors can be required

JWD1000min series OPM provide the FC, SC, ST, LC and other kind of changeable optical fiber adaptors, can adapt for all kinds of the need of the optical fiber connection.

- User-self Calibration Function

Any wavelength can be calibrated by user independently (up to 9 wavelengths calibration can be supported). The device do not need to send back to the original factory for recalibration, it can be deal with in the local Measurement Bureau for the annually calibration.

- Attenuation Chips can be changeable

JWD1000min series OPM support 10,100 and 1000 absorbability attenuation chips, for the convenience of users’ variable attenuation requirements.

- WIN32 API Support

JWD1000 min series OPM provide various API ports based on WIN32, so that the user can develop its own software according to its actual requirements, which includes:

- Data Reading

Read the current wavelength power value (also can customize the relative log, absolute log, and linearity), also can read the detector temperature

2)  Calibration

Calibrate the current wavelength, zero calibration, and adjust the compensation temperature

3)  Setting

Setting the wavelength, Setting the measurements mode (manual, Auto), Setting the ranger, Setting the filter speed, Setting Units, Setting the reference value

New in 2008

JWD1000 Detector Type Min Series Optical Power Meter

Product Specifications

| Model | JWD1000 mini Series OPM |
| --- | --- |
| Wavelength (nm) | 400～1100(Si); 800~1600(Ge) |
| Pre-determined calibrated wavelength (nm) | 850/980/1300/1310/1490/1550/1625（Ge Detector）；660/780/820/850（Si Detector） |
| Self-calibration wavelength qty | 9 Max |
| Uncertainty（Full rang） | 5％ |
| Dynamic Rang（dB） | 70 |
| Unit | Linear：nw、uw、mw、w Log：dB、dBm |
| Displaying speed（Data average） | SLOW：960ms（16 times on average） MEDIUM：480ms（8 times on average） FAST：60ms（uneven） |
| Resolution | 0.001 |
| Communication Port | Standard RS232 Port or USB port |
| Cable length | 1.5m（Standard）or Customized |
| Operation temperature (℃) | ＋10~＋40 |
| Storage temperature(℃) | - 20～＋60 |
| AC adaptor | 9V DC/RS232 or supply by USB cable |
| Warm-up time | 20Min. |
| Weight(kg) | <200g |
| Dimensions(mm) | Ǿ45×40 |

Ordering Information

| Model | JWD1000-A-B-C-D-E mini Power Meter |
| --- | --- |
| Detector Type（A） | - 3mm Ge -60~0dBm - 5mm Ge -60~0dBm - 5mm Si -60~0dBm - 10mm Si -60~0dBm |
| Communication port （B） | - RS232 - USB |
| Cable Length（C） | - 2m（Standard） - 5m (Optional) |
| Attenuator type（D） | - 10 time absorbability attenuation chips (+10~-60dBm) - 100 time absorbability attenuation chips (+20~-50dBm) - 1000 time absorbability attenuation chips (+30~-40dBm) |
| Adaptor Type（E） | - FC - SC - ST - LC - Bare Fiber |

Remark：1、Each attenuation chip is unique relative to the detector.

2、Extra calibration service can be optional on customer requirements

Standard Packages

| Model | include |
| --- | --- |
| JWD1000 mini series Optical Power Meter | Main Unit, RS232 communication cable, FC Adaptor, A piece of Attenuator, AC adaptor, Software CD, Cotton Tampon, User manual. |

New in 2008

JW3201N Bench-top optical Power Meter

JW3201N Bench-top Optical Power Meter is a high precision and wide measurement range test instrument which designed specially aim at passive components factories, R&D institutions and universities. With features of accurate measurements, durable use and easy operation, JW3201 N Bench-top Optical Power Meter becomes a perfect test equipment in fiber optic works which can instead the advanced imported products with its high cost-performance ratio.

JW3201N1 Horizontal detector

JW3201N2 Vertical detector

JW3201N External Detectors Introduction (Refer to right pictures)

- JW3201N1 Horizontal detector:

With 2.0mm Photosensitive area, 0dB InGaAs

- JW3201N2 Vertical detector:

With 3.0*4.0mm Photosensitive area, 0dB TO detector

- The only tip for bare fiber can be optional

- Different external detector can be used on customer requirements

The only Tip for bare fiber

JW-430 bare fiber tip have a V sharp slot, the bare fiber can pull in and out and orientation repeatedly. And it can protect the fiber against the pressure from outside, to avoid fiber tortuosity that will cause the polarization dependent loss. JW-430 can fully encircle the bare fiber, to against the surrounding light come into the fiber, to make the test result more preciously.

V Sharp Slot

N1 type detector or N2 type detector can be matched with JW-430 bare fiber tip

Product Features

- External detector makes the testing operation more convenient.

- 850~1700nm wide wavelength range, wavelength can be adjustable in 1nm increments.

- High accuracy, the uncertainty is±3％

- Display resolution 0.01dB

- Wide Power Measurement range  (+5～-75dBm)

- User self-calibration function

- “Ref” setting enables the user to retrieve and display the measurement that has been stored as a reference. Link loss test result obtained automatically without any manual calculation.

- Adopt the new Data Acquisition technology to ensure optimum signal-noise ratio.

- Use the bare fiber tip to realize power measurement with bare fiber

- Enables data transfer to a PC via RS232 communication port as well as data storage automatically.

New in 2008

JW3201N Bench-top optical Power Meter

Through communication with RS232 port and PC, can realize the optical power monitoring without people, as well as displaying the data on the instrument and the PC at the same time

PC Displaying Interface

Instruments Displaying

Interface

Product Specifications

| Type | JW3201N |
| --- | --- |
| Wavelength Range(nm) | 850～1700 (in 1nm increments) |
| Calibrated Wavelength(nm) | 850，980，1300，1310，1480，1490，1550，1625 |
| Photo Detector | In GaAs |
| Measurement Range (dBm) | +5～-75dBm (other measurement range can be customized) |
| Intrinsic Uncertainty | ±3% |
| Communication interface | RS232 |
| Resolution(dB) | 0.01dB |
| Optical Connector | FC |
| Power Supply | AC90～260V，50Hz |
| Operating Temperature (℃) | -10~+40 |
| Storage Temperature (℃) | -25~+70 |
| Dimension(mm) | 260*190*120 |
| Weight (kg) | 2.75 |

Ordering Information

| Model | Description |
| --- | --- |
| JW3201N－X | 1. JW3201N 1 Horizontal detector 2. JW3201N 2 Vertical detector |
| JW-430 | JW-430 Bare Fiber Tip |

Standard Packages

| Model | Standard Package List |
| --- | --- |
| JW3201N1 Bench-top Power Meter | Main Unit,JW3201N1 horizontal detector, RS232 data transfer cable, Power supply Adaptor, Software CD, Cotton Tampon, User Manual |
| JW3201N2 Bench-top Power Meter | Main Unit, JW3201N1 horizontal detector , RS232 data transfer cable, Power supply Adaptor, Software CD, Cotton Tampon, User Manual |

## Bench-top Optical High Stable Light Source---------JW3102

Joinwit has designed the JW3102 Benchtop highly stable light source

specially used in manufacturing of passive optical components and in testing of lab research. Relying on the high accuracy automatic power control (APC) and the automatic temperature control (ATC) techniques, the JW3102 series light sources has very stable output power and wavelength. JW3102 can deliver 1 to 2 wavelengths upon customer request and laser types can be optional to satisfy customer’s various requests. It is an ideal test equipment for required conditions of widely varying attenuation, highly stable output power and wavelength.

### Features

- Optical Output Power can be adjustable, adjustment range can be custom-made according to the needs

- High stability and reliability

- High accurate ATC and APC control electrocircuit

- Many kinds of laser wavelengths, Customer order is available

- OLED display is employed to make a comfortable seeing.

- USB communication port

## Applications

Manufacturing and Testing of Optical Passive Components

WDM Testing

Lab Testing

Testing of Stability and Linearity

Test instrument calibration

Communication System Test

#### Specifications

| TYPE | JW3102-A(Single wavelength) | JW3102-B(Dua-wavelength) |
| --- | --- | --- |
| Operating Wavelength(nm) | 850,980,1310,1480,1550,1625 | 1310/1550,1310/1480,1480/1550,980/1550 |
| Optical Connector | FC/APC(Other type adapters can be cusomized) |  |
| -3dB Spectral Width(nm) | <0.1 |  |
| -20dB Spectral Width(nm) | <0.5 |  |
| Side mode suppression ratio(dB) | >35(DFB) |  |
| Optical Output Power(mw) | 1~10 |  |
| Output power stability in 15 minutes(dB) | ±0.005 |  |
| Output power stabiltiy in 8 hours(dB) | ±0.05 |  |
| Internal Modulation Frequency | 270Hz,1KHz,2KHz pulsating wave, Dutyfactor 50±10% |  |
| external Modulation Frequency | Digitals Signals(frequency:DC~10MHz,TTL) |  |
| TEC stability | ±0.1℃ |  |
| Operating Voltage | 85~264VAC |  |
| Power Consumption | <15W |  |
| Operating Temperature | -10~+60℃ |  |
| Storage Temperature | -25~+70℃ |  |
| Dimensions(mm) | 300*260*120 |  |
| Net Weight(kg) | 3 |  |
| Remark: All specifications stated here are depend on DFB laser, Specifications like emitter types, central wavelength, output power on demand. |  |  |

##### Standard Packages

| MODEL | INCLUDES |
| --- | --- |
| JW3102 all models | JW3102 Optical Light Source,Power Suppy Cord,Fuze, Instruction Manul and Cotton tampon. |

## ASE Broadband Light Source--------JW3107

JW3107 ASE (Amplified Spontaneous Emission) broadband light source

is designed specially for manufacturing of passive optical components and for testing in lab research. The main body of the light source is composed of high-powered pump laser and erbium doped fiber. Inimitable ATC and APC circuits are used to ensure the stability of the output power. The output power can be adjustable in a certain range by adjusting the APC. The JW3107 broadband ASE light source can operate in the C band, the L band as well as in the C+L band.

### Features

- Hight output power, the max. power can be 100 mw.

- Wide operating bandwidth, covering C-band and/or L-band

- Excellent flatness within spectrum range

- Intelligent micro processing system and long distance control

- High stability and reliability

- Highly accurate ATC and APC control electrocircuit

## Applications

Manufacturing and Testing of Passive Optical Components

WDM Testing

Manufacturing and Testing of EDFA

Lab Testing and Fiber Optic Sensor Systems

Other Optical Fiber Engineering Applications

#### Specifications

| MODEL | JW3107 |
| --- | --- |
| Wavelength Band | C+L Band |
| Operating wavlength | 1525~1610nm |
| Output power | >17dBm |
| Spectrium Density Stability | ±0.05dB/15min |
| Output power stability in 15 minutes | ±0.02dB/15min |
| Output power stabiltiy in 8 hours | ±0.05dB/8hour |
| Flatness | 10dB |
| Output return loss | >45dB |
| TEC stability | ±0.1℃ |
| TEC operating range | 25±5℃ |
| Operating Voltage | 85~264VAC |
| Power Consumption | <15W |
| Operating Temperature | 0~45℃ |
| Storage Temperature | -40~+80℃ |

##### Standard Packages

| MODEL | INCLUDES |
| --- | --- |
| JW3107 all models | JW3107 ASE Broadband Light Source,Power Suppy Cord,Fuze, Instruction Manul and Cotton tampon. |

## Insertion Loss and Return Loss Test Station----JW3307A

# JW3307A Insertion Loss/Return Loss Test Station

# is a high performance loss test station that is designed specially for Optical Passive Components production Test and Lab Test. It combines three different working modes as a return loss meter, optical power and loss meter and a stable laser source in one test station. It features:

- High measurement accuracy

- Accurate analysis to wide dynamic ranger and weak signal

- Two LCD displays used, efficiently reduced eye strain of operators

- Leakage design of optical power meter module and light source module, obviously reduced operation procedures.

- Movitable optical connector set design, easy to clean

- USB Port design, enables data trasfer to a PC via USB port .

#### Specifications

| Model | JW3307A |
| --- | --- |
| Optical Return Loss Test |  |
| Wavelength | 1310/1550nm |
| Optic Connector | FC/APC |
| Return Loss measurement Range | 0 ~ 75dB |
| Calibrated wavelength | 850/1300/1310/1550nm |
| Output Stability of laser source | 0.05dB(1 hour@250C) |
| Measurement accuracy | 0.25dB |
| Resolution | 0.01dB |
| Optical Power and Loss Test |  |
| Wavelength Range | 800~1700nm |
| Calibrated wavelength | 850/1300/1310/1550nm, more other wavelengths can be optional |
| Optic Connector（Power Meter） | Interchangeable FC/SC/ST/2.5mm Universal /1.25mm adaptors |
| Photo detector | InGaAs |
| Display modes | dBm/dB/xW |
| Measurement range | +3 ~ -80 dBm |
| Resolution | Non-linear 0.001dB ；Linear：0.001nw/µW/mW |
| Measurement accuracy | 0.25dB |
| Other Specification |  |
| Communication Port | USB |
| Power Supply | AC 90-260V，50～60Hz |
| Operation Temperature | -5℃～+55℃ |
| Storage Temperature | -25℃～+70℃ |
| Dimensions | 300X260X120mm |
| Weight(kg) | 3 |

#### Standard Packages

| MODEL | INCLUDES |
| --- | --- |
| JW3307A | JW3307A Main Body, Power Supply Cord,FC Adaptor,SC Adaptor, ST Adaptor, 25mm Universal, 1.25mm Universal Adaptor,Instruction Manual,FC/APC-FC/APC patch cord FC/APC-FC/PC patch cord, Cleaning Cotton Swab, Fuse, USB Cable(Optional), PC software (Optional), Foot Peda (Optional) |